Final Report – Segment 3 V&V and Commissioning Flights

# HC-12 – TA Alignment Transfer

## Introduction

This document describes the use of HIPO to transfer the SOFIA TA pupil location and the in-focus TA center-of-flange location, orientation, and scale to the TAAS in the DAOF SIL.

The basic approach for transferring the pupil location is straightforward in principle:

- Determine the location of the secondary mirror in flight using the HIPO pupil viewing optics on the HIPO red side.

- Adjust the position of the Large Chopped Hotplate so its center coincides with the TA pupil position measured in flight.

- Adjust the position of the Small Chopped Hotplate so its center coincides with the TA pupil position measured in flight.

In practice the alignments of the hotplates was complicated by the fact that they are much closer to the instrument in the TAAS than the secondary mirror in the TA is in the actual TA.  This caused focus shifts that had to be worked around.  This is described in section 1.3.

The mapping of the TA focal plane to the TAAS is better understood with some background.  The HIPO red side CCD was mapped to the TA’s SI mounting flange using the HIPO Flange Fiducial jig.  This precisely machined fixture can be mounted in place of the HIPO entrance window and provides a square pattern of pinholes located in the nominal TA focal plane with known spacing, with orientation fixed to the pins on the SI mounting flange to machining tolerances, and with one hole in the pattern located at the center of the flange within machining stack-up tolerances.  The mapping of the HIPO red CCD to this fixture provided the underlying metric for aligning the gyro axes and imagers in the TA to the SI mounting flange.  This activity was carried out during the 2008 line-ops and is described in that report (SCI-US-TRP-SV03-2012, especially the sections on TC-HIPO-01 and TC-HIPO-06) and in Franziska Harms’ thesis.

Because the Flange Fiducial is the primary reference for the TA coordinate systems, transfer of the TA alignment to the TAAS really amounts to verifying the scale and orientation of the TAAS’ FCLS stages using the HIPO red CCD’s known relationship to the Flange Fiducial.  This was done by:

- Determining the position of the FCLS when the focused image is at the known location of the center of the flange.

- Measure a series of X and Y positions with the FCLS to determine the linear relationship between the TAAS FCLS coordinates and the flange coordinate system via the HIPO red CCD’s mapping to the Flange Fiducial.

## Data Acquired

The work for this test was done at the end of the SCAI-2 flight on 20110623 UT and in subsequent work on the TAAS.  Table HC-12-1 lists the relevant data for this test.  Rather than including all the trial images in the process of aligning the hotplates Table HC-12-1 only includes final images showing the as-aligned location of the hotplates.  The hand notes for this test are in the HIPO 2 lab book under these dates.

Table HC-12-1:  HIPO HC-12 Data

| UT Date | UT Range | Red OS | Comments |
| --- | --- | --- | --- |
| SCAI-2 |  |  |  |
| 20110623 | 12:21:19 - 12:23:19 | 279 – 283 | In-flight pupil images, SCAI-2 |
| Post Flight |  |  |  |
| 20110625 | 22:20:10 | 42, 43 | TAAS large chopped hotplate |
|  | 23:06:37 | 51 | TAAS small chopped hotplate |
| 20110626 | 17:01:04 - 18:03:41 | 3-28 | FCLS flange center, stage alignment, and scale |

## Data Analysis

The real-time centroiding capability in LOIS was used to find the in-focus image positions.

The position of the pupil image taken in flight (we used OS279 for this) was found using ds9.  A circular region was created with diameter closely matching the secondary and then moved until it was centered on the secondary as well as possible by eye.  The position uncertainty is less than a pixel.

The main difficulty in this test was to cope with the fact that the hotplates in the TAAS were far too close to the instrument for its pupil viewing lenses to focus.  We tried several approaches for working around this problem.  The following scheme worked.  A small light source was put at the center of the hotplate.  By itself this would not be in focus with the pupil viewing optics, but by adding a small aperture carefully centered in the location of the HIPO entrance window that effectively acted like a pinhole camera it was possible to image the light source on the hotplate.  After determining that our sensitivity to the position of the aperture at the entrance window location was low (less than a pixel for our actual aperture) all that was required was to adjust the hotplate until the image of the light source centered on the hotplate was in the correct position.  The built-in LOIS centroiding capability was used for this.

## Results

### Pupil Position and Hotplates

Figure HC-12-1 shows the in-flight image of the TA pupil with the thin green circular region included.  The circular region is centered at HIPO pixel position (599, 468) and has a diameter of 230 pixels.  The final position of the large chopped hotplate after following the process described in the previous section was (559.6, 468.5).  Taking the hotplate off and re-installing it reproduced this position well, with coordinates (559.0, 468.4).  The small chopped hotplate was treated similarly.  The image of the light source on the hotplate was fainter and the position showed scatter on the level of 2 pixels.  The final position found on OS51 was (598, 471).

Figure HC-12-1.  This shows the in-flight image (HIPO Red OS 279) of the TA pupil with a green circular region described in the text overlaid.

The large chopped hotplate had to be lowered about 18 mm from the position expected by the TAAS engineers in order to reach its aligned state.  The small chopped hotplate needed to be lowered by a comparable amount.  These seemed like very large motions to all of us and left us with a sense of unease.  However, in the fullness of time both FLITECAM and FIFI-LS were aligned using the TAAS hotplates and neither instrument required additional adjustment after being mounted on the TA.

### FCLS

The FCLS alignment results are summarized below:

Focus:

7050 was the best focus position for HIPO.

Flange Center:

Flange fiducial flange center position:  (543.6, 480.9) in HIPO red pixel coordinates.

TAAS (x, z) = (46.081, 31.938) corresponded to (543.2, 480.1) in HIPO red pixels.

This was defined to be the “HIPO Boresight” on the TAAS.

Scale:

The FCLS scale was found to be 12.71 HIPO pixels per mm of FCLS motion.  The HIPO pixel scale is 0.327”/pixel so the effective image scale of the TAAS, referred to the sky, is 4.16 “/mm.

Orientation:

The HIPO red CCD pixel grid was rotated by 0.19 degrees relative to the Flange Fiducial pattern of pinholes (see SCI-US-TRP-SV03-2012, TC-HIPO-01).  The TAAS FCLS axes showed the same orientation error in a vertical scan (i.e. the TAAS was well aligned with the Flange Fiducial) during the alignment transfer activity on 20110626.  A horizontal scan showed a smaller slope but with larger scatter.  This is shown in Figure HC-12-2 below.  FCLS measurements made on 20121005 indicate that the FCLS axes are aligned with the Flange Fiducial with an uncertainty of about 0.05 degrees.

Figure HC-12-2.  A vertical scan in the TAAS on 20110626 (left) showed the same angular offset relative to HIPO pixels that the Flange Fiducial did, suggesting excellent FCLS alignment.  A horizontal scan (right) showed very little slope, but with much larger scatter.